Página del Resumen

- Título de artículo (Español):  Medición de glicemia en hurones: evaluación de tres métodos portátiles de uso humano

- Resumen:

Objetivos: Los equipos portátiles para medición de glucosa en sangre proporcionan resultados poco confiables en especies distintas para las que están calibrados. En el caso de los hurones, los cuadros de hipoglicemia son comunes en la clínica.

Materiales y Métodos: En el presente trabajo se comparó el valor de la glicemia obtenido con tres glucómetros portátiles digitales de uso humano. Se evaluaron 11 muestras sanguíneas de hurones.

Resultados: Los resultados obtenidos sugieren que ninguno de los 3 equipos portátiles evaluados es equivalente en la medición de glucosa en sangre en hurones, al compararlos con los resultados obtenidos en un laboratorio de gabinete, ya que los tres métodos subestiman el valor de la glucosa con respecto al valor obtenido en el laboratorio.

- Conclusiones: Según lo encontrado, los autores sugieren interpretar con cautela los resultados de equipos portátiles para la medición de glicemia en hurones, se sugiere corroborar los resultados con los obtenidos en laboratorio de gabinete.

Palabras Clave: Glucosa, glucómetros portátiles, glicemia, Mustela.

- __________________________________________________________________________

- Título del Artículo (Ingles): Measurement of glucose in ferrets: accuracy of three human portable meters

- Abstract:.

Aims: Point-of-care glucometers had produced unreliable measurements when used in different species to which they were calibrated. In exotics pets clinics, hypoglycemics ferrets are very common.

Methods and Material: Blood samples were obtained from 11 ferrets to compare glycemia values obtained with 3 portable blood glucometers.

Results: The 3 point-of-care glucometers tested underestimated the glycaemia values when compared with the results obtained with the automated biochemical analyzer. Therefore, the results suggest that none of the 3 portable blood glucometers are reliable for measuring blood glucose in ferrets, also indicates an average underestimation in relation to the laboratory analyzer.

Conclusions: The clinicians should be cautious when interpreting glycemia values in ferrets obtained with these 3 point-of-care glucometers, and always corroborate the results with an automated biochemical analyzer

Glucose, point-of-care glucometers, glycemia, Mustela